[Grab your reader’s attention with a great quote from the document or use this space to emphasize a key point. To place this text box anywhere on the page, just drag it.]

[Grab your reader’s attention with a great quote from the document or use this space to emphasize a key point. To place this text box anywhere on the page, just drag it.]

[Grab your reader’s attention with a great quote from the document or use this space to emphasize a key point. To place this text box anywhere on the page, just drag it.]

navify® POC Operations Application

Instrument Driver for use with Abbott i-STAT 1
User Information

August 2025 / 1.9

# Information about the document

| Revision Index | Manual Version | Revision date | Changes |
| --- | --- | --- | --- |
|  | 1.0 | December 2015 | First release |
|  | 1.1 | February 2016 | Update Abbott i-STAT/DE compatibility information |
|  | 1.2 | June 2018 | Update Abbott i-STAT/DE compatibility information |
|  | 1.3 | April 2019 | Update Patient test results section, Update list of predefined tests, test material and QC materials sections |
|  | 1.4 | January 2020 | Update application name, Update Intended Use statement, Update Abbott i-STAT/DE compatibility information, Updates to reflect changes made on the GUI |
|  | 1.5 | June 2020 | Update Abbott i-STAT 1 instrument alarms with an extended list of predefined instrument alarms for the Abbott i-STAT 1 handheld. Update Patient List section Added note about Patient list not sent to Abbott i-STAT/DE when Consilia workflow is active |
|  | 1.6 | April 2021 | Update “General driver information” chapter to include Electronic Simulator results. Update “General connection information” chapter to refer to Abbott i-STAT/DE 2.10 as the minimum required version. Add “Electronic simulator test results”chapter. Update “References” chapter to include test, test material and QC material to be used with Electronic Simulator test results. |
|  | 1.7 | February 2024 | Update the “Electronic simulator test results” chapter. Add “Calibration/Verification (Linearity) test results” chapter. Update branding/product name |
|  | 1.8 | February 2025 | Update the references by adding the High-Sensitivity Troponin I Test in the Abbott i-STAT 1 Tests, Abbott i-STAT 1 test material, Abbott i-STAT 1 QC material, and Abbott i-STAT 1 Linearity material |
|  | 1.9 Table 1 | August 2025 Revision Index | Update for navify POC Operations version 3.3. Remove the limitation that patient list is not supported with Consilia |
| Edition notice | This manual is for users of the navify® POC Operations application. Every effort has been made to ensure that the information contained in this manual is accurate at the time of printing. Not all functionality described in this manual may be available to all users. Roche Diagnostics International Ltd reserves the right to make any further required changes to the software without prior notice. Such changes may not immediately be reflected in this document. Software updates are done by Roche Service representatives. |  |  |
| Intended use | navify® POC Operations is a data management software system providing governance and connectivity to Point of Care devices centralizing the management of professional Point of Care testing service. navify® POC Operations supports workflow configuration of connected Point of Care devices, management of device related data, operators, test orders and patient demographic data including results. These data are automatically received, stored and evaluated based on user configurable criteria and sent to the customers Health Information Systems. The professional personnel can review and audit these data and perform other data management functions. This product is not intended for the diagnosis, screening, monitoring or treatment of patients. |  |  |

| Copyrights | © 2001-2025, Roche Diagnostics International Ltd. All rights reserved. |
| --- | --- |
| Trademarks | Recognized trademarks: - ROCHE, ROCHE OMNI, ACCU-CHEK, ACCU-CHEK INFORM, ROCHE CARDIAC, COBAS B, COBAS, COAGUCHECK and LIFE NEEDS ANSWERS are trademarks of Roche. - Roche and the Roche symbol are registered trademarks of the Roche group. - All other trademarks are trademarks of their owners. |
| Feedback | Every effort has been made to ensure that this guide fulfills its intended use. All feedback on any aspect of this guide is welcome and will be considered during updates. Please contact your Roche representative, should you have any such feedback. |

# Contact addresses

| Manufacturer |  | Roche Diagnostics GmbH Sandhofer Strasse 116 68305 Mannheim Germany Made in Switzerland |
| --- | --- | --- |
| US Distribution |  | Roche Diagnostics Corporation 9115 PO Box 50457 Indianapolis, IN 46250-0457 USA |

Table of contents

Information about the document	2

Contact addresses	5

Abbott i-STAT 1	7

System safety information	8

Patient test results	9

QC test results	10

Liquid QC test results	10

Electronic simulator test results	11

Calibration Verification (Linearity) test results	12

Proficiency test results	13

Configuration	14

Device Configuration	14

Abbott i-STAT/DE Customization Workspace	14

Instruments status	15

Operator List	17

Patient List	18

Reports	20

i-STAT 1 Quality Check Codes (By Location and User)	20

References	22

Abbott i-STAT 1 tests	22

Abbott i-STAT 1 test materials	23

Abbott i-STAT 1 QC materials	24

Abbott i-STAT 1 Linearity materials	24

Abbott i-STAT 1 instrument alarms	24

# Abbott i-STAT 1

| General information and recommendations |  | The Abbott i-STAT 1 handheld is a Point of Care (POC) device, used within a hospital or clinic. Data generated by the Abbott i-STAT 1 handheld can be managed using the navify® POC Operations application. |
| --- | --- | --- |
| General driver information |  | The navify® POC Operations application employs a driver for the Abbott i-STAT 1 handheld. It supports the following functions: - Receiving patient test results from the Abbott i-STAT 1 handheld. - Receiving QC test results from the Abbott i-STAT 1 handheld. - Receiving electronic simulator test results from the Abbott i-STAT 1 handheld. - Receiving Calibration verification (Linearity) test results from the Abbott i-STAT 1 handheld. - Receiving proficiency test results from the Abbott i-STAT 1 handheld. - Receiving quality check codes from the Abbott i-STAT 1 handheld. - Receiving the JAMS/CLEW version from the Abbott i-STAT 1 handheld. - Sending user data from the navify® POC Operations application to Abbott i-STAT 1 handheld. - Sending patient data from the navify® POC Operations application to Abbott i-STAT 1 handheld. |
|  |  | The navify® POC Operations instrument driver for use with the Abbott i-STAT 1 handheld does not work if operating in RiliBÄK quality control mode. |
| General connection information | The navify® POC Operations instrument driver for use with the Abbott i-STAT 1 handheld employs a bi-directional communication with Abbott i-STAT/DE version 2.10 or higher by means of a non-encrypted HTTP connection. From time to time Abbott releases new cartridges and tests that might not be supported by the application. See Abbott i-STAT 1 tests for a list of supported tests for the Abbott i-STAT 1 handheld. See Abbott i-STAT 1 QC materials for a list of predefined QC materials for the Abbott i-STAT 1 handheld. See Abbott i-STAT 1 Linearity materials for a list of predefined QC materials for the Abbott i-STAT 1 handheld. See Abbott i-STAT 1 test materials for a list of predefined test materials for the Abbott i-STAT 1 handheld. |  |
| Abbott i-STAT 1 in navify® POC |  | Instrument type in the navify® POC Operations application is ‘Abbott i-STAT 1’. Instrument type ID in the navify® POC Operations application is ‘IST1’. |

# System safety information

|  | Failure to observe the following safety information may result in incorrect results, data corruption and data losses. |
| --- | --- |
|  | Unprotected HTTP connections between the navify® POC Operations application and Abbott i-STAT/DE may be used to gain unauthorized access to data and for data tampering. Roche recommends the following precaution: - Protect the hospital network against malicious software and hacker attacks. |

# Patient test results

|  | Patient tests can be performed using the Abbott i-STAT 1 handheld. Test results are then recorded and retrievable in the navify® POC Operations application. Along with the patient test results, the following data is being sent: - Patient ID - User ID - Serial number of the Abbott i-STAT 1 handheld - JAMS/CLEW version installed on the Abbott i-STAT 1 handheld - Cartridge / Test lot information - Test measurement date and time - Test result value - Instrument alarms - STATNotes entered on the Abbot i-STAT 1 handheld - Comment entered on the Abbott i-STAT 1 handheld |
| --- | --- |
| Supported Tests | See Abbott i-STAT 1 tests for a list of supported tests for the Abbott i-STAT 1 handheld. |
| Test materials | See Abbott i-STAT 1 test materials for a list of predefined test materials for the Abbott i-STAT 1 handheld. |
| Instrument alarms | See Abbott i-STAT 1 instrument alarms for a list of predefined instrument alarms for the Abbott i-STAT 1 handheld. |
|  | Lots for the predefined test materials are transmitted by the instrument and thus automatically created in the navify® POC Operations application. Lots that are automatically created in the navify® POC Operations application have the expiration date of 31-Dec-9999. The user must manually replace the default expiration date with the correct expiration date printed on the box and/or pouch of the cartridge. |

# QC test results

## Liquid QC test results

|  | Liquid QC tests can be performed using the Abbott i-STAT 1 handheld. Test results are then recorded and retrievable in the navify® POC Operations application. Along with the liquid QC test results, the following data is being sent: - User ID - Serial number of the Abbott i-STAT 1 handheld - JAMS/CLEW version installed on the Abbott i-STAT 1 handheld - QC lot number, level and QC ranges. - Cartridge / Test lot information - Test measurement date and time - Test result value - Instrument alarms - Comment entered on the Abbott i-STAT 1 handheld |
| --- | --- |
|  | Control ranges are transmitted by the instrument to the navify® POC Operations application only if the Pass/Fail determination is automatically performed via an electronic value assignment sheet (eVAS). See the Abbott i-STAT 1 System Manual for instructions on configuring automatic Pass/Fail determination. In case Pass/Fail determination is not configured, control ranges must be manually entered in the navify® POC Operations application. |
| Supported Tests | See Abbott i-STAT 1 tests for a list of supported tests for the Abbott i-STAT 1 handheld. |
| QC materials | See Abbott i-STAT 1 QC materials for a list of predefined QC materials for the Abbott i-STAT 1 handheld. |
| Test materials | See Abbott i-STAT 1 test materials for a list of predefined test materials for the Abbott i-STAT 1 handheld. |
| Instrument alarms | See Abbott i-STAT 1 instrument alarms for a list of predefined instrument alarms for the Abbott i-STAT 1 handheld. |
|  | Lots for the predefined test materials are transmitted by the instrument and thus automatically created in the navify® POC Operations application. Lots that are automatically created in the navify® POC Operations application have the expiration date of 31-Dec-9999. The user must manually replace the default expiration date with the correct expiration date printed on the box and/or pouch of the cartridge. |
|  | Lots for the predefined QC materials are transmitted by the instrument and thus automatically created in the navify® POC Operations application. Lots that are automatically created in the navify® POC Operations application have an expiration date of 31-Dec-9999. The user must manually replace the default expiration date with the correct expiration date printed on the box and/or pouch of the cartridge. |
|  | Liquid QC tests performed with Non-APOC fluids are not supported. |

## Electronic simulator test results

|  | Electronic simulator tests, internal and external, can be performed using the Abbott i-STAT 1 handheld. Test results are then recorded and retrievable in the navify® POC Operations application. Along with the test results, the following data is being sent: - User ID - Serial number of the Abbott i-STAT 1 handheld - JAMS/CLEW version installed on the Abbott i-STAT 1 handheld - Test measurement date and time - Test result value - Comment entered on the Abbott i-STAT 1 handheld - Extra data items such as simulator ID and actual readings taken during the check |
| --- | --- |
| Supported Tests | See Abbott i-STAT 1 tests for a list of supported tests for the Abbott i-STAT 1 handheld. |
| QC materials | See Abbott i-STAT 1 QC materials for a list of predefined QC materials for the Abbott i-STAT 1 handheld. |
|  | By default, the application provides QC lot “DEF-IST1-ESIM-QC” to be used with electronic simulator tests and all electronic simulator test results are associated with this QC lot. |
| Test materials | See Abbott i-STAT 1 test materials for a list of predefined test materials for the Abbott i-STAT 1 handheld. |
|  | By default, the application provides test lot “Internal” to be used with electronic simulator tests and all internal electronic simulator test results are associated with this test lot. For external electronic simulator results, the electronic simulator ID will be used as test lot number. |

## Calibration Verification (Linearity) test results

|  | Calibration verification tests can be performed using the Abbott i-STAT 1 handheld. Test results are then recorded and retrievable in the navify® POC Operations application. Along with calibration verification test results, the following data is being sent: - User ID - Serial number of the Abbott i-STAT 1 handheld - JAMS/CLEW version installed on the Abbott i-STAT 1 handheld - Cal/Ver Kit Number / Linearity lot number. - Cal/Ver Qc Level - Cal/Ver Qc Ranges - Cartridge / Test lot information - Analyte identifier as instrument test name - Measurement unit - Date and time the linearity result was taken - Linearity result value - Instrument alarms - Comment entered on the Abbott i-STAT 1 handheld |
| --- | --- |
|  | Linearity lot number and linearity ranges must exist prior to sending results and have to be manually entered in the navify® POC Operations application. The information about linearity ranges is provided by Abbott. |
| Supported Tests | See Abbott i-STAT 1 tests for a list of supported tests for the Abbott i-STAT 1 handheld. |
| Linearity materials | See Abbott i-STAT 1 Linearity materials for a list of predefined Linearity materials for the Abbott i-STAT 1 handheld. |
| Test materials | See Abbott i-STAT 1 test materials for a list of predefined test materials for the Abbott i-STAT 1 handheld. |
| Instrument alarms | See Abbott i-STAT 1 instrument alarms for a list of predefined instrument alarms for the Abbott i-STAT 1 handheld. |
|  | The test lot is transmitted by the instrument. A test lot needs to already exist in the navify® POC Operations application otherwise the ranges of a specific linearity lot cannot be entered. |
|  | The linearity lot is transmitted by the instrument; the lot must exist and have ranges (must be created prior doing Cal/Ver results on the Abbott i-STAT 1 handheld.) in the navify® POC Operations application. |

# Proficiency test results

|  | Proficiency tests can be performed using the Abbott i-STAT 1 handheld. Test results are then recorded and retrievable in the navify® POC Operations application. Along with the proficiency test results, the following data is being sent: - Proficiency ID - User ID - Serial number of the Abbott i-STAT 1 handheld - JAMS/CLEW version installed on the Abbott i-STAT 1 handheld - Cartridge / Test lot information - Test measurement date and time - Test result value - Instrument alarms - Comment entered on the Abbott i-STAT 1 handheld |
| --- | --- |
| Supported Tests | See Abbott i-STAT 1 tests for a list of supported tests for the Abbott i-STAT 1 handheld. |
| Test materials | See Abbott i-STAT 1 test materials for a list of predefined test materials for the Abbott i-STAT 1 handheld. |
| Instrument alarms | See Abbott i-STAT 1 instrument alarms for a list of predefined instrument alarms for the Abbott i-STAT 1 handheld. |
|  | Lots for the predefined test materials are transmitted by the instrument and thus automatically created in the navify® POC Operations application. Lots that are automatically created in the navify® POC Operations application have the expiration date of 31-Dec-9999. The user must manually replace the default expiration date with the correct expiration date printed on the box and/or pouch of the cartridge. |

# Configuration

|  | Configuration > Devices > Configuration Use this component to create a new instrument configuration or to maintain information for existing instrument configurations. There is a predefined configuration ‘Def. IST1 conf.’ for Abbott i-STAT 1 handheld. |
| --- | --- |
|  | Configure instruments to always use the default configuration for Abbott i-STAT 1 handhelds. Using other configurations might lead to a result mismatch. |

## Device Configuration

|  | Configuration > Devices > Configuration > Device Configuration Use this component to configure Abbott i-STAT 1 specific settings and to access the Abbott i-STAT/DE Customization Workspace. |
| --- | --- |
| Patients > Patient ID Mapping | Used to set which patient ID (i.e. ID1, ID2, or ID3) or visit ID is the source ID to be used in the upload of the patient list to Abbott i-STAT/DE and to which patient ID or visit ID results received from the instrument will be associated. This setting is applied to the navify® POC Operations application only. There is no instrument setting for this. The default value for this field is Patient ID 1. |

## Abbott i-STAT/DE Customization Workspace

|  | The Abbott i-STAT/DE Customization Workspace is a web application provided by Abbott that is used to specify the customizable parameters for Abbott i-STAT 1 handhelds. |
| --- | --- |
|  | To access Abbott i-STAT/DE Customization Workspace |
|  | - Navigate to Configuration > Devices > Configuration component. - Select the default instrument configuration for the Abbott i-STAT 1. - Click Device Configuration. The Device Configuration screen appears. - Click Customization Workspace. The Abbott i-STAT/DE Customization Workspace window appears. |
|  | The information shown in the Abbott i-STAT/DE Customization Workspace is site-specific. |
| Troubleshooting | Institution does not exist |
|  | This error message occurs if there isn’t an Abbott i-STAT 1 handheld assigned to a location of the current site. Use the Device Management > Add new device component to define an Abbott i-STAT 1 instrument that is assigned to a location of the current site. |

# Instruments status

|  | Device Management > By type > Abbott i-STAT 1 Use this component to view JAMS/CLEW versions and the connection status of connected Abbott i-STAT 1 handhelds. |
| --- | --- |
| Instruments grid | Column; Description. Connection; Always ‘Unavailable’. The Abbott i-STAT 1 handheld is not directly connected to navify® POC Operations application, but to Abbott i-STAT/DE. As a consequence, the navify® POC Operations application generally is not aware whether an Abbott i-STAT 1 handheld has been synchronizing at some point during the past 15 minutes.. Last Sync; Date and time when the Abbott i-STAT 1 handheld has last sent a test result to Abbott i-STAT/DE.. Firmware; Displays the JAMS/CLEW version used by the Abbott i-STAT 1 handheld.. |

# Operator List

|  | Using the Operator List feature allows minimizing Operator ID entry errors and prevents uncertified users from running tests. |
| --- | --- |
|  | By default, the feature is turned off. |
|  | The Operator List feature can be enabled via the Abbott i-STAT/DE Customization Workspace. |
|  | When the feature has been enabled, the navify® POC Operations application will send users who have a valid certification for the Abbott i-STAT 1 handheld and are assigned to the same site as the handheld to Abbott i-STAT/DE. Abbott i-STAT/DE will send the list to the Abbott i-STAT 1 handheld when the handheld is docked. |
|  | The navify® POC Operations application sends an up-to-date operator list to Abbott i-STAT/DE every 10 minutes. |
|  | The maximum number of operators that can be uploaded to the Abbott i-STAT 1 handheld is 4000 per site. Exceeding this limit in the navify® POC Operations application can lead to operators being unable to run tests, as only the first 4000 operators are uploaded to Abbott i-STAT/DE. |
|  | The maximum number of characters currently supported by the Abbott i-STAT 1 handheld for the Operator ID field is 15 characters. An operator whose Operator ID defined in the navify® POC Operations application exceeds this maximum is not able to run tests. |

# Patient List

|  | Using the Patient List feature allows minimizing Patient ID entry errors. |
| --- | --- |
|  | By default, the feature is turned off. |
|  | The Patient List feature can be enabled via the Abbott i-STAT/DE Customization Workspace. |
|  | When the feature has been enabled, the navify® POC Operations application will send patients with an admitted visit in the same site as the Abbott i-STAT 1 handheld to Abbott i-STAT/DE. Abbott i-STAT/DE will send the list to the Abbott i-STAT 1 handheld when the handheld is docked. |
|  | The navify® POC Operations application sends an up-to-date patient list to Abbott i-STAT/DE list every 10 minutes. |
|  | The maximum number of patients that can be uploaded to the Abbott i-STAT 1 handheld is 6000 per site. Exceeding this number of admitted patients can lead to operators being unable to run patient tests for some patients. |
|  | If the number of admitted patients exceeds the maximum number of patients supported by the Abbott i-STAT 1 handheld, then navify® POC Operations application will decide which patients are sent to the Abbott i-STAT 1 handheld based on the patient class of admitted visits. The patients having an admitted visit whose patient class is 'Inpatient' are sent with higher priority than patients who only have an admitted visit whose patient class is 'Obstetrics', 'Emergency', 'Preadmit', 'Recurring patient' or 'Auto-generated visits'. The patients who only have an admitted visit whose patient class is 'Outpatient' are sent with the lowest priority. |
|  | The maximum number of characters currently supported by the Abbott i-STAT 1 handheld for the Patient ID field is 15. A patient whose Patient ID defined in the navify® POC Operations application exceeds this maximum will not be uploaded to the Abbott i-STAT 1 handheld. |
|  | The maximum number of characters currently supported by the Abbott i-STAT/DE for the instrument Location Name field is 20. navify® POC Operations application concatenates site and location name, there can be up to 45 characters. The information sent to the Abbott i-STAT/DE will be truncated. |
|  | As the Abbott i-STAT/DE does not support characters outside the ASCII code 32 to 126 range, these characters (in e.g. patient name, middle name, last name instrument location name) are converted into the "?" before sending to Abbott i-STAT/DE. The "?" symbol can be observed in Abbott i-STAT/DE. |

# Reports

|  | Reports > Other reports > Report Generation Use this component to generate Abbott i-STAT 1 specific reports. |
| --- | --- |
| Available reports | - i-STAT 1 Quality Check Codes (By Location and User)i-STAT 1 Quality Check Codes (By Location and User) |

## i-STAT 1 Quality Check Codes (By Location and User)

|  | This report provides a summary of the occurrence and type of Quality Check Codes by user, grouped by location. |  |
| --- | --- | --- |
|  | This report is available in Microsoft Excel format only. |  |
| Filter criteria | Date Range From | The starting date (inclusive) of the quality check codes to be included in the report. |
|  | Date Range To | The ending date (inclusive) of the quality check codes to be included in the report. |
|  | Location ID | The name of the organizational location for which the summary is being generated. |
|  | User Name | The name of the user for which the summary is being generated. |
| Report fields | Site Name | The name of the organizational site to which the handheld was assigned when the quality check code occurred. |
|  | Location Name | The name of the organizational location to which the handheld was assigned when the quality check code occurred. |
|  | User ID | The ID of the user that was using the handheld when the quality check code occurred. |
|  | User Name | The first and last name of the user that was using the handheld when the quality check code occurred. |
|  | Total Cartridges | The total number of patient samples and quality check codes in the specified date range for the operator in the location. |
|  | Total Quality Check | The total number of quality check codes in the specified date range for the operator in the location. |
|  | Quality Check / Cartridge % | The ratio of ‘Total Quality Check’ to ‘Total Cartridges’ (in %). |
|  | Quality Check Codes | The total number of quality check codes per error type in the specified date range for the operator in the location. |
|  | Subtotals | The total number of quality check codes per location. |
|  | Grand total | The total number of quality check codes. |

# References

## Abbott i-STAT 1 tests

The predefined tests for the Abbott i-STAT 1 handheld are:

| Test ID | Test abbreviation | Instrument test name |
| --- | --- | --- |
| 10200 | Cl iS1 | CL |
| 10201 | K iS1 | K |
| 10202 | Na iS1 | NA |
| 10203 | Urea iS1 | UREA |
| 10204 | BUN iS1 | BUN |
| 10205 | iCa iS1 | ICA |
| 10206 | pH iS1 | PH |
| 10207 | pHst iS1 | PHST |
| 10208 | pHt iS1 | PHPT |
| 10209 | Glu iS1 | GLU |
| 10210 | AnGap iS1 | ANGAP |
| 10211 | PO2st iS1 | PO2 |
| 10212 | PO2t iS1 | PO2PT |
| 10213 | PCO2 iS1 | PCO2 |
| 10214 | PCO2st iS1 | PCO2ST |
| 10215 | PCO2t iS1 | PCO2PT |
| 10216 | BE iS1 | BE |
| 10217 | HCO3 iS1 | HCO3 |
| 10218 | sO2 iS1 | SO2 |
| 10219 | Hb iS1 | HB |
| 10220 | Hct iS1 | HCT |
| 10221 | Lac iS1 | LAC |
| 10222 | Crea iS1 | CREA |
| 10223 | U/Cre iS1 | UREACREA |
| 10224 | B/Cre iS1 | BUNCREA |
| 10225 | ACT WBT iS1 | ACTWBT |
| 10226 | aPTT iS1 | APTTWBT |
| 10227 | aPTT-PET iS1 | APTTPET |
| 10228 | aPTT-Rat iS1 | APTTRATIO |
| 10229 | PT iS1 | PTWBT |
| 10230 | PT-PET iS1 | PTPET |
| 10231 | PT-Ratio iS1 | PTRATIO |
| 10232 | PT-INR iS1 | PTINR |
| 10233 | cTnI iS1 | CTNI |
| 10234 | CK-MB iS1 | CKMB |
| 10235 | BNP iS1 | BNP |
| 10236 | TCO2 iS1 | TCO2 |
| 10237 | BhCG iS1 | BHCG |
| 10238 | hCG iS1 | HCG |
| 10239 | TSH iS1 | TSH |
| 10240 | PT+ iS1 | PTPLUS |
| 10241 | PT-INR iS1 | PTINRPLUS |
| 10242 | E-Sim iS1 | ESIM |
| 10243 | hs-TnI iS1 | HS-TNI |
| Table 4 | Abbott i-STAT 1 tests |  |

## Abbott i-STAT 1 test materials

The predefined test materials for the Abbott i-STAT 1 handheld are:

| ID | Description |
| --- | --- |
| 6+T | i-STAT 6+ cartridge |
| ACT-C-T | i-STAT Celite ACT cartridge |
| ACT-K-T | i-STAT Kaolin ACT cartridge |
| aPTT-T | i-STAT aPTT cartridge |
| BhCG-T | i-STAT B-hCG cartridge |
| BNP-T | i-STAT BNP cartridge |
| CG4+-T | i-STAT CG4+ cartridge |
| CG8+-T | i-STAT CG8+ cartridge |
| CHEM8+-T | i-STAT CHEM8+ cartridge |
| CK-MB-T | i-STAT CK-MB cartridge |
| Crea-T | i-STAT Crea cartridge |
| cTnI-T | i-STAT cTnI cartridge |
| E3+-T | i-STAT E3+ cartridge |
| E3+Cl-T | i-STAT E3+Cl cartridge |
| EC4+-T | i-STAT EC4+ cartridge |
| EC6+-T | i-STAT EC6+ cartridge |
| EC8+-T | i-STAT EC8+ cartridge |
| EG6+-T | i-STAT EG6+ cartridge |
| EG7+-T | i-STAT EG7+ cartridge |
| G-T | i-STAT G cartridge |
| G3+-T | i-STAT G3+ cartridge |
| PT-T | i-STAT PT/INR cartridge |
| PTplus-T | i-STAT PTplus cartridge |
| TSH-T | i-STAT TSH cartridge |
| E-Sim-T | i-STAT Electronic Simulator |
| hs–Tni-T | i-STAT hs-TnI cartridge |
| Table 5 | Abbott i-STAT 1 Test materials |

## Abbott i-STAT 1 QC materials

The predefined QC materials for the Abbott i-STAT 1 handheld are:

| ID | Description |
| --- | --- |
| ACT-QC | i-STAT ACT controls |
| B-hCG-QC | i-STAT B-hCG controls |
| BNP-QC | i-STAT BNP controls |
| CHEM8+-QC | i-STAT CHEM8+ controls |
| CK-MB-QC | i-STAT CK-MB controls |
| cTnI-QC | i-STAT cTnI controls |
| i-STAT-QC | i-STAT controls |
| PT-QC | i-STAT PT/INR controls |
| PTplus-QC | i-STAT PTplus/aPTT controls |
| RNAMH-QC | i-STAT RNA Medical Hematocrit controls |
| TriControls-QC | i-STAT TriControls controls |
| E-Sim-QC | i-STAT Electronic Simulator |
| hs-TnI-QC | i-STAT hs-TnI controls |
| Table 6 | Abbott i-STAT 1 QC materials |

## Abbott i-STAT 1 Linearity materials

The predefined Linearity materials for the Abbott i-STAT 1 handheld are:

| ID | Description |
| --- | --- |
| B-hCG-Lin | i-STAT B-hCG calibration verification |
| BNP-Lin | i-STAT BNP calibration verification |
| CHEM8+-1BLin | i-STAT CHEM8+ calibration verification |
| CK-MB-Lin | i-STAT CK-MB calibration verification |
| cTnI-Lin | i-STAT cTnI calibration verification |
| i-STAT-Lin | i-STAT controls calibration verification |
| TriControls-Lin | i-STAT TriControls calibration certification |
| hs-TnI-Lin | i-STAT hs-TnI calibration verification |
| Table 7 | Abbott i-STAT 1 Linearity materials |

## Abbott i-STAT 1 instrument alarms

The predefined instrument alarms for the Abbott i-STAT 1 handheld are:

| Alarm code | Instrument alarm | Description |  | Default severity |
| --- | --- | --- | --- | --- |
| PFBM | <P | Patient result below instrument measurement range |  | Error |
| PFAM | >P | Patient result above instrument measurement range |  | Error |
| CFBM | <C | Control result below instrument measurement range |  | Error |
| CFAM | >C | Control result above instrument measurement range |  | Error |
| FRU | <***> | Result status unknown |  | Error |
| LART | 16 | Low action range triggered |  | Information |
| HART | 32 | High action range triggered |  | Information |
| COVE | 64 | Cutoff value exceeded |  | Information |
| Table 8 | Abbott i-STAT 1 instrument alarms |  |